# ANNEX II + III:	TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

# Contract title: TD 03 Supply of ambulance vehicle for transporting patients	p 1 /…

# Publication reference: 03- 392

Columns 1-2 should be completed by the contracting authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the contracting authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words ‘compliant’ or ‘yes’ are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

LOT 1: Supply of Medical Equipment

| 1. Item number | 2. Specifications required |  | 3. Specifications offered |  | 4. Notes, remarks, ref to documentation |  | 5. Evaluation committee’s notes |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | ## Ambulance Vehicle for Emergency Medical Services ### General Requirements The offered vehicle shall be a new ambulance/sanitary vehicle intended for emergency medical transport and healthcare services, fully equipped and compliant with applicable EU standards and national regulations for ambulance vehicles. Equivalent models from different manufacturers shall be accepted provided that they comply with the minimum technical and functional requirements defined below. # Technical Characteristics of the Base Vehicle Requirement; Minimum Requirement. Vehicle type; Van type ambulance vehicle (L2H2 or equivalent category). Engine type; Diesel engine. Emission standard; EURO 5 or newer. Engine displacement; Minimum 2,000 cm³ with allowed deviation ±10%. Maximum power; Minimum 96 kW (130 HP) with allowed deviation ±10%. Transmission; Manual gearbox with minimum 6 forward gears or automatic equivalent. Maximum speed; Minimum 140 km/h. Fuel tank capacity; Minimum 80 litres. Drive type; Front-wheel drive or equivalent. Maximum allowed mass; Minimum 3,200 kg. Payload capacity; Minimum 1,400 kg. Driving category; Suitable for Category B driving licence. Number of seats in front cabin; 1+2 seats. |  |  |  |  |  |  |
| 1. Item number | 2. Specifications required | 3. Specifications offered |  | 4. Notes, remarks, ref to documentation |  | 5. Evaluation committee’s notes |  |
| 1 continue | # Vehicle Dimensions Equivalent dimensions shall be accepted within the allowed tolerances. Requirement; Minimum Requirement. Overall vehicle length; Minimum 5,400 mm ±10%. Overall vehicle width; Minimum 1,950 mm ±10%. Overall vehicle height; Minimum 2,500 mm ±10%. Wheelbase; Minimum 3,000 mm ±10%. Internal patient compartment length; Minimum 2,950 mm ±10%. Internal patient compartment width; Minimum 1,650 mm ±10%. Internal patient compartment height; Minimum 1,850 mm ±10%. # Driver Cabin Equipment The vehicle must include at minimum: - Air conditioning system with cooling and heating; - ABS + EBD braking system; - ESP/VDC or equivalent electronic stability system; - Power steering with height adjustment; - Driver and passenger airbags; - Electrically operated front windows; - Electrically adjustable side mirrors; - Central locking with remote control; - On-board computer; - Bluetooth/USB radio system or equivalent; - Adjustable driver seat with armrest; - Adjustable steering wheel; - Parking sensors or parking assistance system; - Daytime running lights; - Fog lights; - Safety belts for all passengers; - Spare wheel with same dimensions as original wheels; - Spare wheel tools and jack; - Tinted or athermic windows. Equivalent systems from different manufacturers shall be accepted. |  |  |  |  |  |  |

| 1. Item number | ## 2. Specifications required | 3. Specifications offered | 4. Notes, remarks, ref to documentation | 5. Evaluation committee’s notes |
| --- | --- | --- | --- | --- |
| 1 Continue | # Patient Compartment and Medical Conversion The patient compartment must be professionally converted and suitable for emergency medical transport. Minimum requirements: - Thermal and sound insulation; - Interior walls and ceiling made of ABS, polyester or equivalent disinfectant-resistant material; - Anti-slip and disinfectant-resistant floor made of epoxy resin or equivalent material; - Physical partition between driver and patient compartment with communication sliding window; - Minimum two seats for medical personnel with three-point safety belts; - Cabinets and shelves for medical equipment; - Mounting belts and fixation systems for medical devices; - LED lighting in patient compartment; - Independent air-conditioning system in patient compartment with minimum capacity of 5 kW; - Independent heating system in patient compartment with minimum capacity of 2 kW; - Roof-mounted turbo ventilation/fan system; - Minimum two 12V sockets; - Minimum three 220V sockets; - Inverter minimum 2 kW; - Fire extinguisher minimum 6 kg; - Entry and exit assistance handles; - IV holders; - Control panel accessible to medical staff for lighting, ventilation, heating and air conditioning control. |  |  |  |
| 1. Item number | # 2. Specifications required | 3. Specifications offered | 4. Notes, remarks, ref to documentation | 5. Evaluation committee’s notes |
| 1 Continue | # Mandatory Medical and Transport Equipment The offered ambulance vehicle must include the following equipment: Equipment; Minimum Requirement. Stretcher platform; Installed stretcher fixation platform/system. Main stretcher; Automatic or semi-automatic ambulance stretcher with mattress. Cardio chair; Cardio chair W4 with wheels or equivalent. Stretcher board; Spinal Board for emergency patient transport. Oxygen system; Oxygen cylinder with regulator and breathing mask. Patient fixation system; Safety fixation system for patient transport. Equivalent certified medical transport equipment shall be accepted. # Emergency Light and Sound Signalling The vehicle must include: - Blue LED light bar; - Acoustic emergency siren minimum 100W; - Amplifier with minimum 6 functions; - Front blue flashing lights; - Reflective emergency medical markings; - International emergency medical symbols; - Reflective strips on sides and rear; - Possibility for branding with beneficiary and EU logo. All emergency signalling equipment must comply with applicable national regulations. # Warranty and Service Requirements Requirement; Minimum Requirement. General vehicle warranty; Minimum 3 years or 100,000 km. Anti-corrosion warranty; Minimum 5 years. Paint warranty; Minimum 3 years. Spare parts availability; Minimum 10 years. Authorized service support; Available within the country. |  |  |  |

| 1. Item number | # 2. Specifications required | 3. Specifications offered | 4. Notes, remarks, ref to documentation | 5. Evaluation committee’s notes |
| --- | --- | --- | --- | --- |
| 1 Continue | # Documentation The tenderer shall provide: - Technical catalogues/brochures; - Certificates of conformity; - Vehicle homologation documents; - Warranty documentation; - Service documentation; - Registration documentation in accordance with applicable regulations. # Delivery Period Maximum delivery period: - Up to 180 calendar days including medical conversion and installation of medical equipment. # Compliance The offered vehicle and equipment must comply with: - applicable EU standards; - national ambulance vehicle regulations; - road safety requirements; - emergency medical transport standards. Equivalent technical solutions and equipment from different manufacturers shall be accepted provided that they meet the minimum functional and technical requirements. |  |  |  |